Rigol RSA5065N

VNA S11 Calibration Cheat Sheet

Using CK106A Calibration Kit (N-Type, DC to 6.5 GHz)

Document Version: 1.0  |  Date: February 2026

⚠ CRITICAL WARNING: Offset Length Trap Selecting the CK106A kit from the Cal Kits menu does NOT automatically populate the correct offset parameters for each standard. You MUST manually verify that the Offset Length for Open, Short, and Load are each set to 0.0000 mm. Failure to check this will result in incorrect calibration and misleading measurements.

# 1. Prerequisites

- CK106A calibration kit to hand (Open, Short, Load, Through standards)

- Standards and Port 1 connector clean and undamaged

- Instrument warmed up (minimum 30 minutes recommended)

- Know your target frequency range before starting

# 2. Enter VNA Mode and Set Frequency Range

- Press Mode and select VNA

- Verify the Tracking Generator (TG) is enabled — it must be on for calibration. TG enables by default in VNA mode.

- Press FREQ and set your Start Frequency and Stop Frequency for the band you want to measure.

💡 TIP: Set the frequency range BEFORE calibrating. The calibration is only valid for the frequency range configured at the time of calibration. If you change the range afterwards, you must recalibrate.

- Press Meas and select S11

# 3. S11 Measure Setup — Port Extension

- Go to Meas > S11 Meas Setup > Port Extension

- Set Port Extension to OFF (for local port calibration)

# 4. Select Calibration Kit and Verify Offsets

- Navigate to Meas > Align > Cal Kits > Selected Kit

- Select CK106A

⚠ CRITICAL: Manually Verify Each Standard's Offset Selecting the CK106A kit may NOT auto-populate the offsets. You must check each one individually:

- In Cal Kits, go to Parts Type and select Open — verify Offset Length = 0.0000 mm

- Change Parts Type to Short — verify Offset Length = 0.0000 mm

- Change Parts Type to Load — verify Offset Length = 0.0000 mm

CK106A Correct Offset Values:

| Standard | Parts Type | Offset Length |
| --- | --- | --- |
| Open | Open | 0.0000 mm |
| Short | Short | 0.0000 mm |
| Load | Load | 0.0000 mm |

# 5. Set Amplitude Scale

- Connect the Short standard to Port 1

- Press AMPT (Amplitude) and select All Auto to let the instrument set an appropriate scale based on the Short response

# 6. Run the OSL Calibration

Navigate to Meas > Align. Connect each standard in turn:

- Connect the OPEN standard to Port 1 → Press Open → Wait for sweep to complete

- Connect the SHORT standard to Port 1 → Press Short → Wait for sweep to complete

- Connect the 50Ω LOAD standard to Port 1 → Press Load → Wait for sweep to complete

- Press Done to save the calibration coefficients

⚠ You MUST press Done after all three standards. If you skip this step, the calibration coefficients will not be saved and your measurements will be uncalibrated.

# 7. Verify Calibration

After pressing Done, verify the calibration is working correctly:

## 7a. Load Test (50Ω)

With the Load standard still connected:

- Smith Chart: dot at dead centre (50Ω)

- Return Loss (Log Mag): better than -40 dB

- SWR: very close to 1.0

## 7b. Open Test (Disconnect)

Remove all standards from Port 1 (leave it open):

- Smith Chart: should jump to far right edge (infinite impedance)

- SWR: should read very high / infinity

- Return Loss (Log Mag): close to 0 dB

## 7c. Short Test

Connect the Short standard:

- Smith Chart: should be at far left edge (0Ω)

- Must NOT appear capacitive (below the centre line). If it does, the offset parameters are wrong — go back to Step 4.

💡 TIP: If the measurement does not change when you disconnect a cable or standard, the calibration has not taken effect. Re-run the OSL procedure from Step 6.

# 8. Important Notes

- Calibration is S11 only. Switching to S21 or DTF requires recalibration.

- Recalibrate for each frequency band. A narrow-band cal gives much better accuracy than a wideband sweep.

- Cable measurements: If measuring antenna + cable as a system, calibrate at the instrument port. This measures what the modem/radio actually sees. If measuring just the antenna, calibrate at the antenna end of the cable.

- Port Extension: Use port extension to mathematically shift the reference plane to the far end of a cable without physically calibrating there. Set Port 1 Length to your cable length and Velocity Factor to match your cable type (e.g. 0.66 for RG213).

- Connector care: Always finger-tighten only. Never over-torque N-type connectors on the instrument port.

# Quick Reference: Menu Navigation

| Action | Menu Path |
| --- | --- |
| Enter VNA Mode | Mode > VNA |
| Set Frequency | FREQ > Start Freq / Stop Freq |
| Select Measurement | Meas > S11 |
| Port Extension | Meas > S11 Meas Setup > Port Extension |
| Select Cal Kit | Meas > Align > Cal Kits > Selected Kit |
| Check Offsets | Meas > Align > Cal Kits > Parts Type |
| Run Calibration | Meas > Align > Open / Short / Load > Done |
| Auto Scale | AMPT > All Auto |
| Set Trace Format | Trace > Trace Format > Smith / Log Mag / SWR |